CSI-SSFP Manual

Bruker Paravison 6

For additional information regarding this tool, kindly reach out to the authors at:

- elton.montrazi@gmail.com

- lucio.frydman@weizmann.ac.il

- https://github.com/montrazi

- https://www.weizmann.ac.il/chembiophys/Frydman_group/software

When creating a new study, keeping the “HEAD_PRONE” position is important for the reconstruction pipeline. If you make changes, you'll need to verify the value of “imDataParams.PrecessionIsClockwise” in the reconstruction pipeline and ensure the image orientations are correct.

Having a shorter “Echo Time” is beneficial, which you can achieve by reducing the “Encoding Duration”.

“RF Phase Advance” is set to 180 for alternating flip angles. This logic was adapted from the FISP Bruker sequence.

Both the “Repetition Time” and “Flip Angle Excitation” are configured as optimization choices.

“Repetition Time” is determined by the “Acquisition Duration”. We're employing a “Bandwidth in Hz” of 5 kHz, which is a trade-off considering the time lost due to the digital filter. It's essential to ensure that the number of points is ALWAYS an even number due to a limitation in Bruker's system.

We work with an 8-weight average as it acquires one k-space every 30 seconds (Image matrix 32x32). Moreover, we make use of the “Repetitions” parameter for standard averaging. Feel free to test it out.

Set “Read Offset”, “Phase Offset”, and “Slice Offset” to zero. I haven't yet tested the effects of these offsets; it's possible they may not interfere, but further investigation is needed. While we primarily use it for single slices, it functions for 3D as well. I've previously tested it in a 3D setup while maintaining zero offsets. In the case of a single slice, it performs well for thicker slices, with a recommended thickness of 20 mm. However, slices less than 4 mm didn't yield satisfactory results in vivo.

The carrier frequency needs to be configured to maximize the amplitude of all species and prevent stopbands.

For low concentrations of 2H, adjust the "Receiver Gain" to its maximum value.

This is the final adjustment; modifying other parameters will reset the digital filter to its default setting. It's crucial to set the 'firmware used for digital filtering' to 'DSP_user_defined', as this provides the superior digital filter for SSFP.