ME-SSFP and Weighted ME-SSFP Manual

Bruker Paravison 6

For additional information regarding this tool, kindly reach out to the authors at:

- elton.montrazi@gmail.com

- lucio.frydman@weizmann.ac.il

- https://github.com/montrazi

- https://www.weizmann.ac.il/chembiophys/Frydman_group/software

When creating a new study, keeping the “HEAD_PRONE” position is important for the reconstruction pipeline. If you make changes, you'll need to verify the value of “imDataParams.PrecessionIsClockwise” in the reconstruction pipeline and ensure the image orientations are correct.

“echo number” must be an odd value.

“Echo Time” is set to half of the Repetition Time. To verify the specific echo times, refer to the "DelayArrayForTE" within the "Single Parameter" section. Or refer to the "delay for separate echo" in the "Routine" section.

Try setting a high value for “max grad for rephase readout”, such as 200, to maximize this parameter. Note: Occasionally, this parameter is adjusted when transitioning from the “System” setting of 1H to 2H. This adjustment can influence the previously set “Repetition Time”.

No preparation is necessary for low concentrations of 2H.

“Echo time” is determined by both “Bandwidth” and the number of acquisition points.

The carrier frequency needs to be configured to maximize the amplitude of all species and prevent stopbands.

For low concentrations of 2H, adjust the "Receiver Gain" to its maximum value.

Set “Read Offset”, “Phase Offset”, and “Slice Offset” to zero. I haven't yet tested the effects of these offsets; it's possible they may not interfere, but further investigation is needed. While we primarily use it for single slices, it functions for 3D as well. I've previously tested it in a 3D setup while maintaining zero offsets. In the case of a single slice, it performs well for thicker slices, with a recommended thickness of 20 mm. However, slices less than 4 mm didn't yield satisfactory results in vivo.

Modifying the digital filter to "DSP_imaging_solids" within the "Single Parameter" section will result in an image with a more consistent noise level. Keep in mind that altering other parameters will reset the digital filter to its default state. It's advisable to make this adjustment as the final step.

For the weighted ME-SSFP sequence: Keep "Averages" set to 1 and configure the "weighted average" according to the desired weighting for the encoding dimension. Utilize the "Repetitions" parameter for standard averaging.