## Slide 1: Spatial-Temporal Plane-Wave Image Formation with Sparse Uniform ReSampling (SPURS)
T. Chernyakova, A. Aberdam, E. Bar-Ilan, Y. C. Eldar
Electrical Engineering Department, Technion, Israel
Notes: Hello, my name is Aviad Aberdam.
I will present the work on ‘Spatial-Temporal Plane-Wave Image Formation with Sparse Uniform ReSampling‘, that was done at SAPML lab at the Technion.

## Slide 2: Introduction – Delay-and-Sum
The aperture data is obtained at once
Image is formed line by line
The aperture data is dynamically delayed prior to summation
DAS
Notes: After transmitting a plane-wave the received aperture data contains the information on the entire image.
In standard processing these received signals are dynamically delayed and summed to obtain to obtain one image line.

## Slide 3: Introduction – Delay-and-Sum
The aperture data is obtained at once
Image is formed line by line
  multiplications per transmission
DAS
DAS
DAS
Number of lines
Notes: This process is repeated to form all the image lines.
Thus, the number of multiplications per image is on the order of number of elements times the number of samples in axial direction times the number of image lines. The resulting computational load is often prohibitive for real time implementation.
An alternative approach that reduces the amount of computations, is to obtain the image in the frequency domain.

## Slide 4: Introduction – Spatial-Temporal (ST) Processing
Spectral
Remapping
SPURS
2D-FFT
Lu, ‘97-’06
Liu, ’04
Inverse
2D-FFT
  Vs
  – non-uniform neighbors per uniform sample
Same computational cost
Better image quality!
Notes: In a nutshell, in this approach, 2D Fourier transform of the non-delayed aperture data yields the non-uniform samples of the 2D image spectrum.
After appropriate remapping, we get the uniformly sampled image spectrum and can use inverse 2D Fourier transform to recover the entire image.
The computational complexity in this case is proportional to the order of the remapping method, J (optional: namely, the number of non-uniform samples used by remapping algorithm).
In this work we propose to perform the remapping using SPURS - Sparse Uniform ReSampling – algorithm to obtain a better image quality at the same reduced computational cost.

## Slide 5: Outline
ST processing
Spectral remapping
SPURS
Results (Plane-wave Imaging Challenge in Medical UltraSound, IUS 2016 dataset)
Notes: We next take a closer look at ST processing, and at its core step - spectral remapping.
Then, we introduce SPURS. And finally we present the results obtained by applying our method on PICMUS dataset.

## Slide 6: Motivation
Computational complexity
Image quality
Delay & Sum
Spatial-Temporal Processing
+ SPURS
Lu, ‘97-’06
Liu, ’04
Kruizinga, ‘12
Kiperwas, ‘17
Notes: The main drawback of standard delay and sum processing is high computational complexity.
Spatial Temporal processing developed recently by several groups, significantly reduces the amount of computations that is required to obtain an image.
In this work we propose to preform spatial temporal processing using SPURS which allows to improve image quality both in resolution and contrast, while retaining the above reduction in computational cost.

## Slide 7: Spatial-Temporal Processing
?
Nonlinear Transformation
2D-FFT
Non-Uniform Samples
Notes: We now take a closer look at ST method.
In this approach, we compute 2D Fourier transform of the aperture data.
The resulting aperture spectrum is related to the image spectrum by a nonlinear transformation of the coordinates.
Therefore, if the spectrum of the data is computed on the uniform grid, which is the case when we use Fast-Fourier-Transform, we get non uniform samples of the image spectrum.
The question is how to recover the image from these non uniform samples in frequency.

## Slide 8: Spectral Remapping
Uniform Samples
Linear interpolation
  multiplications
  Reduced image quality
NUFFT
  Improved image quality
  multiplications
SPURS
Same computational cost
Better image quality!
2D-FFT
Linear Interpolation
NUFFT
Nonlinear Transformation
of Coordinates
Non-Uniform Samples
Non-Uniform Samples
Uniform Samples
Kruizinga, ‘12
Notes: The straightforward way is to perform spectral remapping using linear interpolation. Only 2M multiplications are needed to compute M uniform frequency samples. However, the image quality is reduced compared to delay and sum, especially at high image depth.
Another approach is to use Non-Uniform-FFT to compute the non-uniform samples of the aperture spectrum. The above correspond to uniform samples of the image spectrum.
The computational load is slightly increased and is proportional to J, which is the number of neighbors being used to compute each non-uniform sample.
This method allows not only to avoid image degradation but to improve it compared to delay and sum.
In our work, we propose to perform the remapping using SPURS which obtains better image quality with the computational cost of Non Uniform FFT.

## Slide 9: SPURS – SParse Uniform ReSampling
  An algorithm for resampling from a non-Cartesian onto a Cartesian grid developed at SAMPL
  Based on generalized sampling theory and sparse solvers
  Improved performance for MRI
Kiperwas, ‘17
Notes: Spurs is an algorithm for resampling a function from its values on an non-uniform (non-Cartesian) grid to a Cartesian grid developed recently at SAMPL.
It is based on the ideas of generalized sampling theory and sparse solvers and was shown to improve the performance of MRI image reconstruction.

## Slide 10: SPURS – Motivation
Finite support signal
Spectrum spanned by sincs
+
+
=
Notes: To understand the ideas behind SPURS, we consider 1 dimensional signal corresponding to axial image line.
Our signal has a finite support in the image space defined by the maximal depth. The spectrum of such a signal is spanned by a set of shifted sinc functions. It can be noticed that the coefficients of sincs correspond to the uniform samples of the signal spectrum.

## Slide 11: SPURS – Motivation (cont.)
?
is not compactly supported
is a full matrix of sinc coefficients
Storing and inverting is intractable
SPURS provides an accurate approximation of d
at low computational cost
Notes: At the output of Spatial Temporal processing we get non uniform samples of the signal spectrum, and we are interested to recover the uniform samples d.
In matrix form, the vector of the uniform samples d is related to the measurement vector b by a matrix A which contains the values of the sinc function.
Thus, to recover d we need to invert matrix A.
But since a is a sinc function which is not compactly supported, the matrix A is a full matrix of sinc coefficients.
The dimensions of matrix A in typical imaging scenario do not allow to store and invert it.
SPURS provides a method to approximate d with high accuracy at low computational cost without a need to use matrix A.

## Slide 12: SPURS
?
  Step 1 – Projection onto spline subspace
  is compactly supported kernel (e.g. B-spline of degree )
  is a sparse matrix (only non-zero entries at each row)
  To find only multiplications are required
  Step 2 – Projection onto sinc subspace
  Orthogonal projection by LSI (Linear Space Invariant) filtering in k-space
Kiperwas, ‘17
Notes: The approximation of d is computed in two steps.
Instead recovering vector d directly, we aim to compute the projection of the signal of interest to a subspace spanned by a compactly supported kernel, for example, B-spline.
This can be done by computing the expansion coefficients c, that are related to our non linear measurements b through matrix Q. Due to compact support of kernel q, the matrix Q is sparse. (For example, if we use B-spline of degree 3, there are only 4 non-zero entries at each row of the matrix)
To find vector c, we formulate a weighted regularized least squares problems. Exploiting sparsity of matrix Q, we can find c with only p+1 times M multiplications.
The second step is to project the resulting approximation to the original subspace spanned by sincs using simple LSI filtering.
This second projection yields uniformly spaced samples of the signal spectrum. From which we can move back to the image space by inverse FFT.

## Slide 13: SPURS
?
Kiperwas, ‘17
  The weighting matrix is introduced to the least squares problem
  Accounts for varying spectral noise density
  The contrast is improved
Notes: The weighting matrix W introduced to the least squares problem allows to account for varying spectral noise density, leading to improved contrast as we will see next.

## Slide 14: Results Simulation – Point Reflectors
Results S
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Notes: We now present the results.
We start with a simulation of point reflectors and compare SPURS to delay and sum and Spatial Temporal processing with both linear interpolation and Non uniform FFT.
Visually, these two images look much cleaner.
To see the differences between NUFFT and SPURS let’s zoom in one of the reflectors. Here we can see that the resolution is comparable but the side lobes are much more prominent in the NUFFT , even though it’s hard to see it on the screen.
Now, Let’s take a look on lateral and axial scanlines.

## Slide 15: Results Simulation – Point Reflectors
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Notes: We can see that in the lateral direction, ST processing gives much better resolution compared to delay and sum. Again we can see that in terms of side lobes SPURS outperforms all the other methods.
In particular, the side-lobes of SPURS are from 10 to 20 dB below those of NUFFT.
The resolution in axial direction is comparable for all the methods, again side-lobes of SPURS are lower compared to NUFFT and linear interpolation.
Thus, we expect that SPURS will outperform the other methods in terms of contrast.

## Slide 16: Results Simulation – Contrast
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Notes: We next take a look at a phantom of anechoic cysts. Again we can see that the performance of linear interpolation degrades with the depth, while both NUFFT and SPURS improve the contrast.
Zoom in shows that the cyst is darker in the image obtained with SPURS.

## Slide 17: Results Simulation Contrast
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Chart: Contrast
| DAS | NUFFT | SPURS
Contrast | 19.9 | 21.667 | 22.133
Notes: It can also be seen in a lateral scanline that the side-lobes of SPURS within the cyst a lower for SPURS.
To verify it quantitatively we computed the contrast based on all the cysts in the image and it can be seen that SPURS provides additional improvement of contrast.

## Slide 18: Results Experiment Resolution
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Notes: We next take a closer look at resolution using experimental data.

## Slide 19: Results Experiment Resolution
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Chart: Lateral Resolution
| DAS | Linear Interpolation | NUFFT | SPURS
Lateral Resolution | 0.89655 | 0.77051 | 0.76496 | 0.76575
Chart: Axial Resolution
| DAS | Linear Interpolation | NUFFT | SPURS
Axial Resolution | 0.57395 | 0.57706 | 0.57499 | 0.57395
Notes: In the lateral scanline corresponding to one of point reflectors we can see that the delay and sum gives much wider mainlobe.
Here we can see that the lateral resolution of the ST methods is the same.
In axial direction the resolution of both NUFFT and linear interpolation is degraded compared to DAS and SPURS.

## Slide 20: Results Experiment Contrast
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Notes: We can see that the above results hold for an experimental scan as well

## Slide 21: Results Experiment Contrast
Delay & Sum
Linear Interpolation
NUFFT
SPURS
Chart: Contrast
| DAS | NUFFT | SPURS
 | 14.4 | 16.65 | 17.2
Notes: And the contrast is improved by SPURS

## Slide 22: Conclusions
Implementation of spectral remapping using SPURS algorithm
Computationally complexity of NUFFT
Superior performance in contrast
?
Nonlinear Transformation
SPURS
Notes: To sum up, we presented a method for implementation of frequency remapping step of ST processing which allows to improve the contrast, while retaining reduced computational cost.

## Slide 23: Thank You